## System Analysis
MIT-E VENT System Design spreadsheet
Spreadsheet to deterministically design ambu bag squeezer using pure rotary linkage with sontinuous rotary motion input. COVID sadness can bring world together to work as a team, and then maybe next step fix climate change, hence one over sadness equals happiness 1/:-( = :-). see www.ooseh.com
created by Prof. Alex Slocum March 21, 2020, see "Revisions" tab
Inputs in BLACK, Outputs in RED | NOTE this is for the gears specified in Bottom Drive Design 10Beta from SA had created from drawings from MIT
Respiration |  |  |  | Comments
Respiration rate
Peak for shorter periods | resprate | 40 | breaths per minute
Average for mechanism life and motor power calculations | resprateavg | 30 | breaths per minute
arc of travel of finger to compress bag | arcfing | 26.5762 | degrees
seconds per breath | spb | 2
Inspiratory/expiratory ratio | ier | 1:4
Inspiration ratio | insp | 1 | portion of cycle
Expiration ratio | expr | 4 | portion of cycle
Total parts | tpb | 5 | total portions
Time to inhale | tinsp | 0.4 | seconds
Time to exhale | texh | 1.6 | seconds
Portion of inhalation to ramp | pcr | 0.1 | percentage
time to ramp to pressure | trtp | 0.04 | seconds
effective average rpm of finger | rpmfing | 11.0734 | rpm
Drive transmission ratio (enter gear sizes below) | TRpf | 1.6
required instantaneous rpm of gearmotor shaft | rpmgmr | 17.7175 | rpm | : =trpf*Rpmfing
Forces on "Hands" from airway pressure in bag
peak airway pressure | pap | 30.7692 | mmHg
|  | 40 | cmH2O
factor higher to force from bag through circuit | papeta | 0.9
bag pressure | pbagmm | 34.188 | mmHg
|  | 0.044984 | atm
|  | 0.66127 | psi
|  | 0.0045611 | N/mm^2
| cmH2O | 44.4444 | cmH2O
Bag diameter | Dbag | 125 | mm
| Lbag | 125 | mm
| Areabag | 15625 | mm^2
force on bag | Fbag | 71.2673 | N
|  | 16.0151 | lbs | force on EACH finger
Estimated Motor Size Required for the System
Life cycle
Breaths per day | bpdavg | 43200
Days of life of machine before rebuild | dolm | 180
Total machine cycles required | tmcd | 7776000 | cycles
Number of oscillations of gearmotor shaft | noms | 12441600 | cycles
Angular oscillation of motor shaft | nodoms | 42.5219 | degrees
Motor Power estimate to start design
Radial velocity of bag sides during inhalation | Vrbagsides | 156.25 | mm/sec | : =Dbag/2/tinsp
Power that must be applied to compress bag during inhalation portion of breathing cycle | Pbaginhale | 22.271 | Watts | : =2*Fbag*Vrbagsides/1000. NOTE the factor of 2 is because one motor drives both fingers
mechanical system effieicny (gearmotor and connection to bag mechanism) | etamech | 0.8
Oversize motor for 100% duty cycle | ovszf | 1
Estimated total peak motor mechanical power to select candidate motors | Pmotorest | 27.8388 | Watts
Geared Hands Design: MIT-E Vent | [5 empty] | 2
Orientation of the pinion-to-drive gear contact force mid-bag compression: for a "pure" top drive Theta = 0. For a "pure" bottom drive drive Theta = 180. For a "pure" side drive Theta = 90
Theta (angle at which the pinion force engages the finger drive gear) | Thetapingear | 180 |  | Original MIT-E Vent bottom drive
distance finger center to finger contact point | R_1 | 139.7 | mm |  |  | #REF!
|  | 5.5 | inch
distance finger center to center plane (finger-to-finger gear radius) | R_2 | 38.1 | mm
|  | 1.5 | inch
diameter of driver gear | Dgear | 76.2 | mm | 48 tooth 16 pitch 14.5 degree pressure angle
From "Gears" worsksheet where number of teeth and Pitch are entered for finger-to-finger gear |  | 76.2
Radius of driver gear | Rgear | 38.1 | mm | For bottom drive driven gear diameter equals 2x distance finger center to center plane
|  | 1.5 | inch
diameter of finger support shaft | Dbear | 9.525 | mm
|  | 0.375 | inch
coefficient of friction of finger support bearing | mu | 0.1
Force on a each "hand" as bag is compressed | Ffinger | 71.2673 | N
Force on gear teeth between fingers to share torque from gearmotor | Fgtg | 261.3135 | N
Required motor pinion tooth force on driven gear to move both "Hands" together to compress bag | Fgear | 522.6271 | N | The pinion must be able to drive BOTH fingers, hence the factor of "2"
|  | 117.4443 | lbf
Torque generated by driven gear which is then shared by both Hands | Tgear | 19912.0914 | N-mm
|  | 176.1664 | lb-in
Hand shaft support Bearing forces
Y direction | Fbx | -593.8944
X direction | Fby | 261.3135
Force on finger support shaft (and bearings) | Fbear | 648.8415 | N
radial load per bearing | Fperbear | 324.4208 | N
|  | 72.9035 | lb
Torque lost to shaft bearing friction | Tlsbf | 309.0108 | N-mm
% torque lost to shaft bearing friction | ptlsf | 1.552% |  |  | 0
Gearmotor
Pinion diameter (NOTE cannot go smaller than 1" for 12 pitch gear) | Dpinion | 47.625 | mm
From "Gears" worsksheet where number of teeth and Pitch are entered |  | 1.875 | inch | 30 tooth 16 pitch 14.5 degree pressure angle
Transmission ratio motor pinion-to-finger | TRpf | 1.6
gearbox output torque needed | Tgm | 12445.0571 | N-mm
|  | 110.104 | lb-in
gearbox output speed needed | gosn | 17.7175 | rpm
NOTE if during the inhale time tinhale (e.g. 0.4 seconds total) if assume x% for ramp up and x% for ramp down , then over total travel of Y degrees, the motor speed required (upper limit) is Y/(tinhale*(1-2x))
Percent of inhale time for ramp up and ramp down | xpramp | 10%
Maximum desired gearbox speed needed to enable ramp up | gosnact | 22.1468 | rpm | : =gosn/(1-2*xpramp)
Motor Selection and operating parameters |  |  |  |  | Motor Data from "Motor Options" workbook
N motor size (Nema 23, 34 or Gearmotor 1, 2,3. see workbook "Motor Options") | Nmsize | 1 |  |  | Motor number | 23 | 34 | 1 | 2 | 3
Stall torque | Tstallsm | 22.5156 | N-mm | : =IF(NMsize=23,G88,IF(NMsize=34,H88,IF(NMsize=1,I88,IF(NMsize=2,J88,K88)))) | Stall torque (N-m) | 1.8014 | 4.2033 | 22.5156 | 44.76 | 28
| Tstallsmm | 22515.621 | N-mm | : =IF(NMsize=23,G88,IF(NMsize=34,H88,IF(NMsize=1,I88,IF(NMsize=2,J88,K88))))
Free speed | wfree34 | 75 | rpm |  | no load speed (rpm) | 900 | 720 | 75 | 28 | 70
max radial force 20mm from face | Fradmaxm | 222.5 | N |  | allowable side load (N) | 75 | 220 | 222.5 | 222.5 | 200
required speed for respiration cycle | rsarc | 17.7175 | rpm
torque available at required speed | NemaTas | 17196.6956 | N-mm
torque required for desired performance | tgm | 12445.0571 | N-mm
|  | 110.104 | lb-in
Number of motors required to provide required torque | nmr | 1
Pinion tooth force at pitch line at motor torque at required speed for required cm H2O | PFPLr | 522.6271 | N | : =tgm/(Dpinion/2)/NMr NOTE that the radial force depends on how many motors being used
|  | 117.4443 | lbf
Percent of maximum shaft load | pomsl | 234.9%
Is radial force on motor shaft too high? |  | force too high? |  | at 35 mm HG load is
cm H2O this torque will provide | cmH2O | 44.4444 | cm H2O
|  | 34.188 | mmHg
Pinion tooth force at pitch line at maximum motor torque at required speed | PFPL | 722.1709 | N
|  | 162.2856 | lbf
Percent of maximum shaft load |  | 324.6%
Is radial force on motor shaft too high? |  | force too high?
cm H2O this torque will provide | cmH2Omax | 80.409 | cm H2O
|  | 61.8531 | mmHg

## gearbox life
For AM PG 71 and 188 gearbox, front bearing face is flush with front face of mounting flange
Bearings
width across bearing outer faces | wabof | 0.455 | in
spacing between bearings | sbb | 0.16 | in
width of bearing | wob | 0.1475 | in
distance between bearings' centerlines | dbbcl | 0.3075 | in
L10 dynamic radial load capacity | drlc | 460 | lbs
million rotation Moment capacity of bearing pair | mcbp | 141.45 | lb-in
Mounting
distance from front face of gearbox to side face of pinion gear | dddgbg | 0.5
width of gear clamped to hex shaft | woac | 0.46 | in
distance from center of pinion to gearbox face | dcagbf | 0.73 | in
distance from center of pinion to center of stiffness | dcacs | 0.9575 | in
radial load on shaft from system | rslgm | 117.4443 | lb
Moment on bearings | mob | 112.4529 | lb-in
Lifetime estimate
Millions of rotations potentially possible with arm close to gearbox face | mrp | 1.9902 | million rotations
Days of life | dol | 46.0695 | days

## Gears
Spur_Gears.xls
Spreadsheet to estimate gear tooth strength
Production gears must be designed using the Lewis Form Factor or FEA
By Alex Slocum 1/18/01, last modified 9/28/2007 by Alex Slocum
Enters numbers in BOLD, Results in RED
Inputs
| Motor Pinion | Finger to Finger
Torque, T (in-lb) | 110.104 | 88.0832
N-mm | 12445.0571 | 9956.0457
Pressure angle, f (deg, rad) | 14.5 | 0.25307
Pitch, P | 16 | 16
arc angle tooth (degrees) | 12 | 7.5
Number of teeth on pinion, Np | 30 | 48
Number of teeth on gear, Ng | 48 | 48
Center distance tolerance, Ctol (inches) | 0.005 | 0.005
Effective face width, w (inches) | 0.25 | 0.25
Pinion material | 420 SS |  | 4140 or 1144 Steel |  | Mild Steel
Pinion yeild stress, sigp (psi) | 60000 | 60000
Gear Material | 420 SS
Gear yield stress, sigg (psi) | 60000 | 60000
Stress concentration factor at tooth root, scf | 1.5 | 1.5
Outputs
Gear ratio, mg | 1.6 | 1 | : =Ng/Np
Pinion pitch diameter, Dp (inches) | 1.875 | 3 | : =Np/P
Gear pich diameter, Dg (inches) | 3 | 3 | : =Ng/P
Center distance, C (inches) | 2.4425 | 3.005 | : =0.5*(Dp+Dg)+Ctol
Tooth thickness, tt (inches) | 0.098175 | 0.098175 | : =0.5*PI()/P
Addendum, a (inches) | 0.0625 | 0.0625 | : =1/P
Dedendum, b (inches) | 0.077 | 0.077 | : =1.2/P+0.002
Clearance, cl (inches) | 0.0145 | 0.0145 | :=0.2/P+0.002
Pinion tooth force (normal to radius), Fp (lbf) | 117.4443 | 58.7221 | :=T/(Dp/2)
Check, should be equal to Fgear, else pinion diameters do not match | YES
Face width assumed to be the molded thickness, force/finger (lbf) | 117.4443 | 58.7221
Gear tooth separation force (parallel to radius) Fsp (lbf) | 30.3732 | 15.1866 | : =Fp*TAN(f)
Total radial load on motor shaft (lbf) | 121.3082 | 60.6541 | :=SQRT(Fsp^2+Fp^2)
Tooth section parameters
Chordal area, Ac (inches^2) | 0.024544 | 0.024544 | : =w*tt
First moment of area, Q (inches^3) | 0.0004016 | 0.0004016 | :=w*tt^2/6
2nd moment of area ("moment of inertia"), I (inches^4) | 1.9713e-05 | 1.9713e-05 | :=w*tt^3/12
Distance Nuetral axis to outer fiber, cc (inches) | 0.049087 | 0.049087 | :=tt/2
| Motor side gear tooth stress |  | Finger-to-Finger gear stress
Tooth stresses | Stress | Safety Factor | Stress | Safety Factor
Shear stress of the tooth (F/A) (psi) | 4785.1107 | 7.2393 | 2392.5553 | 7.2393 | :=Fp/Ac
Bending shear stress (FQ/wI) (psi) | 9570.2214 | 3.6197 | 4785.1107 | 3.6197 | :=Fp*Q/(w*I)
Bending stress (F*b*c/I) (psi) | 33777.3309 | 1.7763 | 16888.6654 | 3.5527 | :=Fp*b*cc/I*scf
Von Mises stress (psi) | 34779.3032 | 1.7252 | 17389.6516 | 3.4503 | :=SQRT(3*B39^2+B41^2)
Assumed line contact stress across facewidth (very generous assumption)
Contact stress (see "geartoothcontact stress" workbook) | 162346.3223 | psi | 116258.6161 | psi
subsurface shear stress | 48703.8967 | psi | 34877.5848 | psi
equivelant tensile stress for life calculation (csets) | 82796.6244 | psi | 59291.8942 | psi
Gear tooth bending Fatigue life
material |  | 420 SS
yield strength | sigyield | 60000 | psi
Pinion | stress ratio | 0.57966 | Von Mises stress at root / yield stress
| pinionfw | 0.25 | inch (contact width with finger gear)
| rotlifepin | 1106978.9236 | cycles
| numtpc | 4 | number of teeth contact per cycle
| pinionlifedays | 64.0613 | days
Finger to finger gear | stress ratio | 0.28983 | Von Mises stress at root / yield stress
| fingfw | 0.25 | inch
| rotlifefing | 3327129278.4587 | cycles
| numtfg | 4 | number of teeth contact per cycle
| fingerlifedasys | 192542.2036 | days
Contact Stress Fatigue Life (pinion limited) |  | pinion | finger-to-finger
Case hardened gear teeth? |  | yes | no
Maximum shear stress/(ultimate tensile/1.7) | pinioncsr | 0.38836 | 0.9882
Life as a function of contact stress | pinionhertzfl | 874560869.2052 | 55.4219 | cycles
| pinionhertzdays | 12652.7904 | 0.00080182 | days
Case hardening
depth of maximum contact shear stress | dcsm | 0.024958 | inch
minimum depth factor | mdf | 2
case depth |  | 0.049916 | inch
RC hardness | rch | 45
equivelant tensile strength at hardness | sighardmpa | 1471.0651 | MPa Slocum Precision Machine Design page 335
| sighardened | 213197.8411 | psi
| sigunhard | 60000 | psi
Life curve for steel core (e.g., scale from https://icme.hpc.msstate.edu/mediawiki/index.php/Fatigue-life_curve-4130_Steel for a good quality steel)
sigmax1 |  | 60000
exp_1 |  | 1
siginfinte |  | 30000
exp_2 |  | 7
slope | mlife | -5000
intercept | blife | 65000
Life curve for steel core (e.g., scale from https://icme.hpc.msstate.edu/mediawiki/index.php/Fatigue-life_curve-4130_Steel for a good quality steel)
sigmax1ch |  | 213197.8411
exp_1ch |  | 1
siginfintech |  | 106598.9206
exp_2ch |  | 7
slope | mlifech | -17766.4868
intercept | blifech | 230964.3279

## geartoothcontact stress
Spreadsheet adopted from FUNdaMENTALS (2.75) to check the line contact stress between motor pinion (steel) and plastic drive gear
Hertz_contact_line.xls
To determine Hertz contact stress between two cylinders
By Alex Slocum, last modified 2/10/2004 by Alex Slocum
Enters numbers in BOLD, Results in RED
| pinion-to-finger | finger to finger
Smaller cylinder 1 diameter, d_1 (mm) | 8.89 | 8.89 | the approximate curvature of the tooth face of a 16 pitch gear as measured from SW gear tooth
Larger cylinder 2 (or flat plane) diameter, d_2 (mm) | 8.89 | 8.89
Length, L (mm) | 6.35 | 6.35 |  | hardened | unhardened
Applied load per roller, F (N) | 509.5614 | 261.3135 | Yield stress allowable (ksi) | 213.1978 | 60
Elastic modulus Eone (N/mm^2) | 200000 | 200000 | Equivelent modulus, Ee (N/mm^2) | 109182.2251
Elastic modulus Etwo (N/mm^2) | 200000 | 200000 | Equivelent modulus, E1e (N/mm^2) | 218364.4503
Poisson's ratio vone | 0.29 | 0.29 | Equivelent modulus, E2e (N/mm^2) | 218364.4503
Poisson's ratio vtwo | 0.29 | 0.29
Hardened teeth | yes | no
Ultimate tensile stress, sigult (N/mm^2) | 1471.0651 | 414
Depth below contact surface for evaluating deflection, do (microns) | 100 | 100
Rectangular contact zone width, 2b (mm) | 0.09121 | 0.065317
Contact pressure, qcyl (N/mm^2) | 1120.1896 | 802.1845 | For d_1 pressed against a rigid flat plate
Deflection motion of d_1 center, defl_1 (mm) | 0.0012787 | 0.00065575 | Conact area half-width b_1 | 0.045605
Deflection motion of d_2 center, defl_2 (mm) | 0.0012787 | 0.00065575 | For d_2 pressed against a rigid flat plate
Total relative displacement of the cylinder's centers, dcyls (mm) | 0.0025574 | 0.0013115 | Conact area half-width b_2 | 0.045605
Stress factor: Must be less than 1
Maximum shear stress/(ultimate tensile/1.7) | 0.38836 | 0.9882
Manufacturing issues
Surface roughness, Ra (mm) | 0.005
Potential induced contact width, Bra (mm) | 0.42166
Compare to sphere-on-flat-plate of same diameter:
Ronemaj | 4.445
Ronemin | 4.445
Rtwomaj | 1000
Rtwomin | 1000
Phi (degrees) | 0
costheta | 0
theta | 1.5708
alpha | 1.003
beta | 1.0014
lambda | 0.74904
Equivelent radius Re | 2.2127
ellipse c | 0.25003
ellipse d | 0.24963
Contact pressure, q | 3898.0785
Stress ratio (must be less than 1) | 2.6498
Deflection at the one contact interface
Deflection (µunits) | 14024.3095
Stiffness (load/µunits) | 0.036334
for circular contact a = c, a | 0.25003
Depth at maximum shear stress/a | 0.63393
Max shear stress/ultimate tensile | 0.8883

## Hex drive
pinion to gearbox shaft
hex shaft size | hsd | 0.5 | in
|  | 12.7 | mm
Width of contact face | wcf | 0.25 | in
|  | 6.35 | mm
torque to be transmitted | tgm | 110.104 | lb-in
constant from the torque integral | Chexpres | 0.063968
Number of hex faces assumed to transmit the torque | Nhftt | 3
contact pressure on hex faces | hcpf | 36719.6086 | psi
|  | 253.274 | MPa
Material yield strength |  | 60000 | psi
expected life |  | 452979.2189 | cycles
|  | 6.5535 | days

## Motor Options
Enter numbers in BLACK, equivelants are calculated in RED, and then these values are transferred to the "System Analysis" spreadsheet
Motor number referenced on "System Analysis" sheet
1 | AndyMark PG71 gearmotor. https://www.andymark.com/products/hex-pg-series-gearboxes-options
| Stall torque | Tstall | 16.6 | lb-ft | 22.5156 | N-m
| Free speed | wfree | 75 | rpm | 1.309 | rad/sec
| estimated allowable side load on shaft | ? No spec found, so estimate | 50 | lb | 222.5 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 7.3682 | W
2 | AndyMark PG188 gearmotor. https://www.andymark.com/products/hex-pg-series-gearboxes-options
| Stall torque | Tstall | 33 | lb-ft | 44.76 | N-m
| Free speed | wfree | 28 | rpm | 0.48869 | rad/sec
| estimated allowable side load on shaft | ? No spec found, so estimate | 50 | lb | 222.5 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 5.4685 | W
3 | High torque windshield wiper motor
| Stall torque | Tstall | 20.6434 | lb-ft | 28 | N-m
| Free speed | wfree | 70 | rpm | 1.2217 | rad/sec
| estimated allowable side load on shaft | ? No spec found, so estimate | 44.9438 | lb | 200 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 8.5521 | W
23 | Nema 23 Stepper motor
| Stall torque | Tstall | 255 | oz-in | 1.8014 | N-mm
| Free speed | wfree | 900 | rpm | 15.708 | rad/sec
| estimated allowable side load on shaft |  | 16.8539 | lb | 75 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 7.0741 | W
34 | Nema 34 Stepper motor
| Stall torque | Tstall | 595 | oz-in | 4.2033 | N-m
| Free speed | wfree | 720 | rpm | 12.5664 | rad/sec
| estimated allowable side load on shaft |  | 49.4382 | lb | 220 | N
| est peak power (1/2 stall torque X 1/2 free speed) |  |  |  | 13.2051 | W
https://szbobet.en.alibaba.com/product/60516551996-803622807/BM_Brand_High_precision_4_Axis_8_5Nm_1204oz_in_Nema_34_Stepper_Motor_with_braker_Driver_Kit_For_CNC_Router.html
https://hardware-cnc.nl/en/shop/stappenmotoren/stepper-motor-34/stepper-motor-nema-34-4,5nm-detail
NEMA 23
https://cnc1.lv/en/stepper-motors/48-nema-23-stepper-motor-23hs9430-with-280-nm.html

## Bearings
Analysis of bearings supporting the finger support shaft
Radial force per bearing from gear drive and hand loads centered | Frbearing | 324.4208 | N
width of bearing including flange | wbearing | 6.35 | mm
0.25
flange width of bearing | flngbear | 1.1684
0.046
diameter bearings | dbearing | 9.525 | mm |  | This is set equal to the finger support shaft on "Systems Analysis" workbook
|  | 0.375 | in
Potential eccentricty of load aplied | ela | 25 | mm
Finger force | Ffinger | 71.2673
twisiting moment on finger | tmof | 1781.6832 | N-mm | : =Fbearing*Ela
spacing of bearings | sob | 50 | mm | narrow if a single center metal gear used with wide array of fingers at the top
force couple applied to bearings | fcob | 35.6337 | N
Worst case net applied radial force to bearing | Fnetbear | 360.0544 | N
|  | 80.9111 | lbf
contact stress on bearings | sigbear | 7.2952 | N/mm^2 | : =Fbearing/(C4*Dbearing)
|  | 1057.6615 | psi
go to igus bearing life calculator to get bearing
https://www.igus-cad.com/?cul=en-US&ArtNr=J350FI-0607-04&mandant=USA
part J350FI-0607-04 - iglide® J350 type F, inch sizes
hours of life | hobl | 12156
| daobl | 506.5

## lung compliance notes
Lung Compliance Update; Alexander Slocum, Jr, MD PhD; on 3/22/2020; sent to Albert Kwon and Christoph Nabzdyk 3/22/2020
Clinical scenario: 300-400 lb patient who has a lung compliance of 3? Also patient is bucking, having difficulty breathing.
Normal lung compliance, lower bound | LCNL | 50 | cmH2O
Normal lung compliance, upper bound | LCNU | 100 | cmH2O
Change from lower bound
Bad compliance, ARDS | LCA | 3 | cmH2O | 16.6667
Resp rate lower bound |  | 8 | breaths
Resp rate upper bound |  | 40 | breaths
PEEP | PEEP | 15 | cmH2O | This exists everywhere in the breathing curcuit AFTER the valve; so inspiratory pressure of 30 is between bag & valve
Inspiratory pressure | Pins | 30 | cmH2O
Driving pressure (at machine) | Pmach | 15 | cmH2O
Driving pressure (at mouth) | Pmouth | 15 | cmH2O
Anticipated required airway pressure (need torque to generate this much press) | Preq | 30 | max

## refs
https://en.wikipedia.org/wiki/Mechanical_ventilation
https://www.sciencedirect.com/topics/medicine-and-dentistry/mean-airway-pressure
https://www.sciencedirect.com/topics/medicine-and-dentistry/airway-pressure

## Revisions
Version | date | who | edit
1 | 2020-03-24 | AHS Sr | added NEMA 34 drve parameters and cm H2O equivelant
2 | 2020-03-24 | AHS Sr | added motor rpm for breathe cycle
3 | 2020-03-24 | AHS Sr | added select motors and do motor calcs, and added details on gear teeth for both pinion and finger to finger
4 | 2020-03-24 | AHS Sr | add in Albert Kwin and Alex Jr breathing edits
5 | 2020-03-25 | AHS Sr | added radial load additional design info
6 | 2020-03-26 | AHS Sr | added motor references
7 | 2020-03-27 | AHS Sr | generalizing for any design
8 | 2020-03-28 | AHS Sr | generalizing for any design and adding fixturing and hex drive calcs
9 | 2020-03-29 | AHS Sr | generalizing for any design and adding fixturing and hex drive calcs
10 | 2020-03-29 | AHS Sr | finger support bearings loaded eccentrically
11 | 2020-03-29 | AHS Sr | motor cycle calcs
12 | 2020-03-29 | AHS Sr | clean up formatting and gave generic name to the spreadsheet (changed from "Top Drive" to "System Design"
13 | 2020-03-29 | AHS Sr | motor power and system design to be based on average bpm, not hyperventilate
14 | 2020-04-01 | AHS Sr | enter specific parameters for MIT Bottom Drive for days of life study of gearbox, bearings, gears
15 | 2020-04-08 | AHS Sr | Life assessment for 40 cm H2O max. Controls team puts on a software limit and a label
1 | 2020-04-09 | AHS Sr | Rename Event Design Spreadsheet with Life assessment 2020.04.09 V1, and add more detailed fatigue life calcs for hardened gears to find viable design
2 | 2020-04-10 | AHS Sr | run analysis for 1/4" driver gear as RC45 and follower annealed as per request. Added calculate tooth contact stress for gears so can select hardened or not

## Bag fixture by Vee
Bag1eD | 50 | mm
Bag2eD | 40 | mm
Veeangle | 45 | degrees
deltaCenter | 3.5355 | mm | distance center heigh shifts when supporting Bags of different end diameters in a Vee

## Expected Tidal Volume
cross section area of bag
before compression | 18.1 | 18.1 | in^2
after compression | 4.82 | 4.82 | in^2
length of compression zone | 3 | 4 | in^3
volume compressed | 39.84 | 53.12 | in^3
tidal volume | 653 | 870 | cm^3
% of ends also compressed | 20% | 20%
expected tital volume | 783 | 1045
actual measured by tests | >1000 cm^3